## Autodrive
Energy | Mode | MU | Accessory | X1 | X2 | Y1 | Y2 | RepRate | Depth mm | N | M1 | M2 | M3
6X | FIXED | 100 |  | 5 | 5 | 5 | 5 | 600 | 100 | 2 | -1.332e-07 | -1.332e-07
15X | FIXED | 100 |  | 5 | 5 | 5 | 5 | 600 | 100 | 2 | -1.513e-07 | -1.513e-07
6X | FIXED | 100 |  | 5 | 5 | 5 | 5 | 600 | 16 | 2 | -1.994e-07 | -1.993e-07
15X | FIXED | 100 |  | 5 | 5 | 5 | 5 | 600 | 27 | 2 | -1.966e-07 | -1.966e-07
6E | FIXED | 100 | A10 |  |  |  |  | 600 | 14.8 | 2 | -2.183e-07 | -2.182e-07
6E | FIXED | 100 | A10 |  |  |  |  | 600 | 21.5 | 2 | -1.584e-07 | -1.59e-07
9E | FIXED | 100 | A10 |  |  |  |  | 600 | 22.1 | 2 | -2.197e-07 | -2.197e-07
9E | FIXED | 100 | A10 |  |  |  |  | 600 | 31.5 | 2 | -1.686e-07 | -1.687e-07
12E | FIXED | 100 | A10 |  |  |  |  | 600 | 30.8 | 2 | -2.24e-07 | -2.241e-07
12E | FIXED | 100 | A10 |  |  |  |  | 600 | 44.5 | 2 | -1.673e-07 | -1.675e-07
15E | FIXED | 100 | A10 |  |  |  |  | 600 | 38.7 | 2 | -2.229e-07 | -2.231e-07
15E | FIXED | 100 | A10 |  |  |  |  | 600 | 55.5 | 2 | -1.698e-07 | -1.701e-07
18E | FIXED | 100 | A10 |  |  |  |  | 600 | 46.9 | 2 | -2.198e-07 | -2.201e-07
18E | FIXED | 100 | A10 |  |  |  |  | 600 | 65.5 | 2 | -1.745e-07 | -1.739e-07

## MEASURE
TG 51 Monthly Calibration ¦ | [5 empty] | 21EX 1833 HSF AR
|  |  |  | Physicist | Rex Cardan, Ph.D. |  | Number of MU Used | 100
|  |  |  | Date | 2017-04-19 |  | Temperature | 20.9 | C
|  |  | Ion Chamber (Cal: 11/09/2016) |  | Extradin A12 (XA023335) - AR |  | Pressure | 749.4 | mmHg
|  |  | Electrometer (Cal: 11/05/2015) |  | SI Max 4000 PLUS - AR |  | Ctp | 1.0104
Photons
6X
| Mraw (nC) |  |  | M (nC) | Dose Conversion |  | Energy Check | Output (cGy/MU) | OK?
Depth (cm) | Read 1 | Read 2 | Average | Mraw*Ppol*Pion*Ctp*Pelec | *Ndw - Co60 | *kQ |  | M(10) / (TMR(10)*#MU)
1.6 | 19.94 | 19.93 | 19.935 | 20.1955 | 97.4835 | 97.0234 | 66.82% |  | PASSED
10 | 13.32 | 13.32 | 13.32 | 13.494 | 65.1357 | 64.8283 | 1.0096 | 1.010 cGy/MU @ dmax | PASSED
15X
| Mraw (nC) |  |  | M (nC) | Dose Conversion |  | Energy Check | Output (cGy/MU) | OK?
Depth (cm) | Read 1 | Read 2 | Average | Mraw*Ppol*Pion*Ctp*Pelec | *Ndw - Co60 | *kQ |  | M(10) / (TMR(10)*#MU)
2.7 | 19.66 | 19.66 | 19.66 | 19.9527 | 96.3115 | 95.8569 | 76.96% |  | PASSED
10 | 15.13 | 15.13 | 15.13 | 15.3552 | 74.1197 | 72.658 | 1.0105 | 1.011 cGy/MU @ dmax | PASSED
Electrons
6E
| Mraw (nC) |  |  | M (nC) | Dose Conversion (kR50) |  |  | Output (cGy/MU) | OK?
Depth (cm) | Read 1 | Read 2 | Average | Mraw*Ppol*Pion*Ctp*Pelec | *Ndw - Co60 | *kecal*k'R50/CF |  | M(dref) /PDD(dref)
1.48 | 21.83 | 21.82 | 21.825 | 22.3804 | 108.0302 | 100.6588 | 1.0074 | 1.007 cGy/MU @ dmax | PASSED
2.15 | 15.84 | 15.9 | 15.87 | 16.2739 | 78.5539 |  | 0.72657 | 0.727 cGy/MU @ dmax | PASSED
9E
| Mraw (nC) |  |  | M (nC) | Dose Conversion (kR50) |  |  | Output (cGy/MU) | OK?
Depth (cm) | Read 1 | Read 2 | Average | Mraw*Ppol*Pion*Ctp*Pelec | *Ndw - Co60 | *kecal*k'R50/CF |  | M(dref) /PDD(dref)
2.21 | 21.97 | 21.97 | 21.97 | 22.4863 | 108.5414 | 100.0223 | 1.0008 | 1.001 cGy/MU @ dmax | PASSED
3.15 | 16.86 | 16.87 | 16.865 | 17.2613 | 83.3205 |  | 0.76718 | 0.767 cGy/MU @ dmax | PASSED
12E
| Mraw (nC) |  |  | M (nC) | Dose Conversion (kR50) |  |  | Output (cGy/MU) | OK?
Depth (cm) | Read 1 | Read 2 | Average | Mraw*Ppol*Pion*Ctp*Pelec | *Ndw - Co60 | *kecal*k'R50/CF |  | M(dref) /PDD(dref)
3.08 | 22.4 | 22.41 | 22.405 | 22.8822 | 110.4524 | 100.9384 | 1.0107 | 1.011 cGy/MU @ dmax | PASSED
4.45 | 16.73 | 16.75 | 16.74 | 17.0965 | 82.525 |  | 0.74618 | 0.746 cGy/MU @ dmax | PASSED
15E
| Mraw (nC) |  |  | M (nC) | Dose Conversion (kR50) |  |  | Output (cGy/MU) | OK?
Depth (cm) | Read 1 | Read 2 | Average | Mraw*Ppol*Pion*Ctp*Pelec | *Ndw - Co60 | *kecal*k'R50/CF |  | M(dref) /PDD(dref)
3.87 | 22.29 | 22.31 | 22.3 | 22.8067 | 110.088 | 100.0645 | 1.0067 | 1.007 cGy/MU @ dmax | PASSED
5.55 | 16.98 | 17.01 | 16.995 | 17.3812 | 83.8989 |  | 0.75753 | 0.758 cGy/MU @ dmax | PASSED
18E
| Mraw (nC) |  |  | M (nC) | Dose Conversion (kR50) |  |  | Output (cGy/MU) | OK?
Depth (cm) | Read 1 | Read 2 | Average | Mraw*Ppol*Pion*Ctp*Pelec | *Ndw - Co60 | *kecal*k'R50/CF |  | M(dref) /PDD(dref)
4.69 | 21.98 | 22.01 | 21.995 | 22.5521 | 108.8589 | 98.5721 | 1.0167 | 1.017 cGy/MU @ dmax | PASSED
6.55 | 17.45 | 17.39 | 17.42 | 17.8612 | 86.2161 |  | 0.81691 | 0.817 cGy/MU @ dmax | PASSED
Note on B4: Rex Cardan: This logo button clears all the previous values (MACRO) to reset the sheet for new values
Note on F12: Rex Cardan: The correction values are dependent on the machine and electrometer selected at the top of the page.
Note on H12: Rex Cardan: The value depends on the ion chamber selected. It is looked up via the "Chambers" tab
Note on I12: Rex Cardan: Rex Cardan: Looking at the current accelerator PDD (10) value from the "Machines" tab. Compares to within 2%
Note on B25: Rex Cardan, Ph.D.: This is the dref value + 0.5*rcav. This worsheet does not include pgrad corrections and instead picks measurement points that have the correction included (effective point of measurement is at dref)
Note on B30: Rex Cardan, Ph.D.: This is the dref value + 0.5*rcav. This worsheet does not include pgrad corrections and instead picks measurement points that have the correction included (effective point of measurement is at dref)
Note on B35: Rex Cardan, Ph.D.: This is the dref value + 0.5*rcav. This worsheet does not include pgrad corrections and instead picks measurement points that have the correction included (effective point of measurement is at dref)
Note on B40: Rex Cardan, Ph.D.: This is the dref value + 0.5*rcav. This worsheet does not include pgrad corrections and instead picks measurement points that have the correction included (effective point of measurement is at dref)
Note on B45: Rex Cardan, Ph.D.: This is the dref value + 0.5*rcav. This worsheet does not include pgrad corrections and instead picks measurement points that have the correction included (effective point of measurement is at dref)

## Chambers
CHAMBERS
Chamber Label | kQ Calc (from Table 1 - TG51) | [5 empty] | %dd(10)x |  | kQ |  |  |  |  | k'R50 |  |  |  |  | Last ADCL Calibration
| 58 | 63 | 66 | 71 | 81 | 93 | Low X | High X | Low X | High X | kecal | rcav (cm) | Ndw-Co60 | 6E | 9E | 12E | 15E | 18E
Extradin A12 (XA023335) - AR | 1 | 0.999 | 0.996 | 0.99 | 0.972 | 0.948 | 66.6 | 76.4 | 0.99528 | 0.98028 | 0.906 | 0.305 | 48270000 | 1.0284 | 1.0171 | 1.0087 | 1.0033 | 0.99945 | 2016-11-09 | 184
PTW TN30013 (1937) - RMC | 1 | 0.996 | 0.992 | 0.984 | 0.967 | 0.946 | 66.6 | 76.4 | 0.99104 | 0.97482 | 0.898 | 0.275 | 54960000 | 1.0284 | 1.0171 | 1.0087 | 1.0033 | 0.99945 | 2012-04-18 | 1850
Extradin A12 (XA132053) - RMC | 1 | 0.999 | 0.996 | 0.99 | 0.972 | 0.948 | 66.6 | 76.4 | 0.99528 | 0.98028 | 0.906 | 0.305 | 49150000 | 1.0284 | 1.0171 | 1.0087 | 1.0033 | 0.99945 | 2015-10-05 | 585
Extradin A12 (XA022324) - HSROC | 1 | 0.999 | 0.996 | 0.99 | 0.972 | 0.948 | 66.6 | 76.4 | 0.99528 | 0.98028 | 0.906 | 0.305 | 48040000 | 1.0284 | 1.0171 | 1.0087 | 1.0033 | 0.99945 | 2015-02-15 | 817
[7 empty] | 66.6 | 76.4 | #NAME? | #NAME? |  |  |  | 1.0284 | 1.0171 | 1.0087 | 1.0033 | 0.99945
[7 empty] | 66.6 | 76.4 | #NAME? | #NAME? |  |  |  | 1.0284 | 1.0171 | 1.0087 | 1.0033 | 0.99945
[7 empty] | 66.6 | 76.4 | #NAME? | #NAME? |  |  |  | 1.0284 | 1.0171 | 1.0087 | 1.0033 | 0.99945
[7 empty] | 66.6 | 76.4 | #NAME? | #NAME? |  |  |  | 1.0284 | 1.0171 | 1.0087 | 1.0033 | 0.99945
[7 empty] | 66.6 | 76.4 | #NAME? | #NAME? |  |  |  | 1.0284 | 1.0171 | 1.0087 | 1.0033 | 0.99945
[7 empty] | 66.6 | 76.4 | #NAME? | #NAME? |  |  |  | 1.0284 | 1.0171 | 1.0087 | 1.0033 | 0.99945
[7 empty] | 66.6 | 76.4 | #NAME? | #NAME? |  |  |  | 1.0284 | 1.0171 | 1.0087 | 1.0033 | 0.99945
[7 empty] | 66.6 | 76.4 | #NAME? | #NAME? |  |  |  | 1.0284 | 1.0171 | 1.0087 | 1.0033 | 0.99945
[7 empty] | 66.6 | 76.4 | #NAME? | #NAME? |  |  |  | 1.0284 | 1.0171 | 1.0087 | 1.0033 | 0.99945
[7 empty] | 66.6 | 76.4 | #NAME? | #NAME? |  |  |  | 1.0284 | 1.0171 | 1.0087 | 1.0033 | 0.99945
[7 empty] | 66.6 | 76.4 | #NAME? | #NAME? |  |  |  | 1.0284 | 1.0171 | 1.0087 | 1.0033 | 0.99945
[7 empty] | 66.6 | 76.4 | #NAME? | #NAME? |  |  |  | 1.0284 | 1.0171 | 1.0087 | 1.0033 | 0.99945
[7 empty] | 66.6 | 76.4 | #NAME? | #NAME? |  |  |  | 1.0284 | 1.0171 | 1.0087 | 1.0033 | 0.99945
[7 empty] | 66.6 | 76.4 | #NAME? | #NAME? |  |  |  | 1.0284 | 1.0171 | 1.0087 | 1.0033 | 0.99945
Note on J2: Rex Cardan, Ph.D.: These values will change based on the machine you choose on the Measure tab.

## Electrometers
The purpose of this tab is to keep a list of electrometers which can be used for the monthly QA. The date of the most recent ADCL calibration is also stored, so as to alert a user on the "Measure" tab when 1) It is a few months away from needed calibration and 2) Its calibration is expired
ELECTROMETERS
Electrometer Label | Electrometer Factor | Last ADCL Calibration
PTW 10005 (T111067¦High Mode) -AR | 1.003 | 2015-05-13 | 730
SI Max 4000 (E021141 ¦ High) - RMC | 0.999 | 2013-05-15 | 1458
SI Max 4000 PLUS - AR | 0.999 | 2015-11-05 | 554

## Machines
MACHINES MACHINES MACHINES
BASIC INFORMATION | [7 empty] | PHOTON INFORMATION | [8 empty] | ELECTRON INFORMATION | [39 empty] | CORRECTION INFORMATION | [18 empty] | SAD Calibrated
| ENERGIES SUPPORTED | [6 empty] | %dd(10)x |  | Clinical PDD(10) |  | Clinical TMR¦PDD(10) |  | ISF | dmax (cm, SSD,10x10) |  | I50 (cm) |  |  |  |  | dref (cm) |  |  |  |  | Electron Depth Dose at dref |  |  |  |  | I50 (cm, A10/A20 [for high E]) |  |  |  |  | I80 (cm, A10/A20 [for high E]) |  |  |  |  | I90 (cm, A10/A20 [for high E]) |  |  |  |  | Energy Tolerances (I80-2mm) |  |  |  |  | Energy Tolerances (I80+2mm) |  |  |  |  | Cone Factor |  |  |  |  | Pion | [6 empty] | Ppol
Machine Label | Low X | High X | E1 | E2 | E3 | E4 | E5 | Low X | High X | Low X | High X | Low X | High X | Inv. Sq. | Low X | High X | E1 | E2 | E3 | E4 | E5 | E1 | E2 | E3 | E4 | E5 | E1 | E2 | E3 | E4 | E5 | E1 | E2 | E3 | E4 | E5 | E1 | E2 | E3 | E4 | E5 | E1 | E2 | E3 | E4 | E5 | E1 | E2 | E3 | E4 | E5 | E1 | E2 | E3 | E4 | E5 | E1 | E2 | E3 | E4 | E5 | Low X | High X | E1 | E2 | E3 | E4 | E5 | Low X | High X | E1 | E2 | E3 | E4 | E5
21EX 1833 HSF AR | 6 | 15 | 6 | 9 | 12 | 15 | 18 | 66.6 | 76.4 | 0.668 | 0.766 | 0.777 | 0.87 | 1.21 | 1.6 | 2.7 | 2.364 | 3.547 | 4.955 | 6.246 | 7.562 | 1.3235 | 2.0539 | 2.9232 | 3.7203 | 4.5328 | 0.9992 | 0.9994 | 0.9987 | 0.994 | 0.9695 | 2.3 | 3.6 | 5 | 6.3 | 7.6 | 2 | 3 | 4.3 | 5.4 | 6.4 | 1.8 | 2.8 | 3.9 | 4.9 | 5.7 | 0.9 | 0.9 | 0.85 | 0.84 | 0.82857 | 0.6 | 0.7 | 0.71429 | 0.73333 | 0.75 | 1 | 1 | 1 | 1 | 1 | 1.0039 | 1.0058 | 1.015 | 1.0138 | 1.0118 | 1.0129 | 1.0151 | 0.99981 | 0.99966 | 1.0009 | 1.0002 | 1 | 1.0003 | 1.0007 | TRUE
iX HSF 5326 AR | 6 | 15 | 6 | 9 | 12 | 15 | 18 | 66.15 | 77.6 | 0.668 | 0.766 | 0.777 | 0.87 | 1.21 | 1.6 | 2.7 | 2.312 | 3.492 | 4.94 | 6.175 | 7.39 | 1.2914 | 2.02 | 2.914 | 3.6764 | 4.4266 | 0.9991 | 0.999 | 0.9984 | 0.9944 | 0.975 | 2.3 | 3.6 | 5 | 6.3 | 7.6 | 2 | 3 | 4.3 | 5.4 | 6.4 | 1.8 | 2.8 | 3.9 | 4.9 | 5.7 | 0.9 | 0.9 | 0.85 | 0.84 | 0.82857 | 0.6 | 0.7 | 0.71429 | 0.73333 | 0.75 | 1 | 1 | 1 | 1 | 1 | 1.0036 | 1.0065 | 1.0138 | 1.0151 | 1.0134 | 1.0149 | 1.0145 | 1.0004 | 1.0005 | 0.99942 | 1.0001 | 1.0003 | 0.99989 | 1.0002 | TRUE
iX 5867 RMC | 6 | 15 | 6 | 9 | 12 | 15 | 18 | 66.1 | 77.91 | 0.661 | 0.766 | 0.769 | 0.877 | 1.21 | 1.5 | 2.7 | 2.396 | 3.63 | 5.076 | 6.438 | 7.845 | 1.3433 | 2.1052 | 2.9979 | 3.8388 | 4.7075 | 0.999 | 0.999 | 1 | 0.994 | 0.974 | 2.37 | 3.65 | 5.18 | 6.54 | 7.798 | 1.957 | 3 | 4.35 | 5.425 | 6.27 | 1.767 | 2.717 | 3.968 | 4.893 | 5.392 | 0.90526 | 0.87067 | 0.85236 | 0.83759 | 0.82278 | 0.65472 | 0.70769 | 0.72771 | 0.74619 | 0.76073 | 1 | 1 | 1 | 1 | 1 | 1.0038 | 1.0063 | 1.0142 | 1.0148 | 1.0141 | 1.0135 | 1.0132 | 1.0011 | 1.001 | 1.0003 | 1.0005 | 1.0006 | 1.0012 | 1.0004 | TRUE
Note on K3: Rex Cardan, Ph.D.: This is used to check the energy ratio
Note on M3: Rex Cardan, Ph.D.: This is the clinical TMR(10,10x10) or PDD (10,10x10) that is used to move the dose from 10cm depth to dmax. If the machine is calibrated for 1cHy/MU at dmax, this value should be the TMR (10), otherwise it should be the PDD(10).
Note on O3: Rex Cardan, Ph.D.: This exists to correct for setup differences between SSD and Isocenter. If it is a TMR calibrated machine with 100 cm SSD setup, the ISF would be (110/100)^2. If it is a SSD calibrated machine, ISF = 1 and PDD(10) should be used in the TMR column.
Note on AV3: Rex Cardan, Ph.D.: These values are interpolated between the R80 and R90 values.
Note on BA3: Rex Cardan, Ph.D.: These values are interpolated between the R80 and R50 values.
Note on BF3: Rex Cardan, Ph.D.: This is in place to adjust the electron measurements in the case the A10 cone is not used. If the A10 is used, just put in 1.0
Note on BK5: Correction measurements taken by Rex Cardan, Ph.D. on 4/14/2014 using the Extradin A12 (XA132053) - RMC chamber.
Note on BL5: Correction measurements taken by Rex Cardan, Ph.D. on 4/14/2014 using the Extradin A12 (XA132053) - RMC chamber.
Note on BM5: Correction measurements taken by Rex Cardan, Ph.D. on 4/14/2014 using the Extradin A12 (XA132053) - RMC chamber.
Note on BN5: Correction measurements taken by Rex Cardan, Ph.D. on 4/14/2014 using the Extradin A12 (XA132053) - RMC chamber.
Note on BO5: Correction measurements taken by Rex Cardan, Ph.D. on 4/14/2014 using the Extradin A12 (XA132053) - RMC chamber.
Note on BP5: Correction measurements taken by Rex Cardan, Ph.D. on 4/14/2014 using the Extradin A12 (XA132053) - RMC chamber.
Note on BQ5: Correction measurements taken by Rex Cardan, Ph.D. on 4/14/2014 using the Extradin A12 (XA132053) - RMC chamber.
Note on BR5: Correction measurements taken by Rex Cardan, Ph.D. on 4/14/2014 using the Extradin A12 (XA132053) - RMC chamber.
Note on BS5: Correction measurements taken by Rex Cardan, Ph.D. on 4/14/2014 using the Extradin A12 (XA132053) - RMC chamber.
Note on BT5: Correction measurements taken by Rex Cardan, Ph.D. on 4/14/2014 using the Extradin A12 (XA132053) - RMC chamber.
Note on BU5: Correction measurements taken by Rex Cardan, Ph.D. on 4/14/2014 using the Extradin A12 (XA132053) - RMC chamber.
Note on BV5: Correction measurements taken by Rex Cardan, Ph.D. on 4/14/2014 using the Extradin A12 (XA132053) - RMC chamber.
Note on BW5: Correction measurements taken by Rex Cardan, Ph.D. on 4/14/2014 using the Extradin A12 (XA132053) - RMC chamber.
Note on BX5: Correction measurements taken by Rex Cardan, Ph.D. on 4/14/2014 using the Extradin A12 (XA132053) - RMC chamber.
Note on BK7: Correction measurements taken by Rex Cardan, Ph.D. on 3/10/2014 using the Extradin A12 (XA132053) - RMC chamber.
Note on BL7: Correction measurements taken by Rex Cardan, Ph.D. on 3/10/2014 using the Extradin A12 (XA132053) - RMC chamber.
Note on BM7: Correction measurements taken by Rex Cardan, Ph.D. on 3/10/2014 using the Extradin A12 (XA132053) - RMC chamber.
Note on BN7: Correction measurements taken by Rex Cardan, Ph.D. on 3/10/2014 using the Extradin A12 (XA132053) - RMC chamber.
Note on BO7: Correction measurements taken by Rex Cardan, Ph.D. on 3/10/2014 using the Extradin A12 (XA132053) - RMC chamber.
Note on BP7: Correction measurements taken by Rex Cardan, Ph.D. on 3/10/2014 using the Extradin A12 (XA132053) - RMC chamber.
Note on BQ7: Correction measurements taken by Rex Cardan, Ph.D. on 3/10/2014 using the Extradin A12 (XA132053) - RMC chamber.
Note on BR7: Correction measurements taken by Rex Cardan, Ph.D. on 3/10/2014 using the Extradin A12 (XA132053) - RMC chamber.
Note on BS7: Correction measurements taken by Rex Cardan, Ph.D. on 3/10/2014 using the Extradin A12 (XA132053) - RMC chamber.
Note on BT7: Correction measurements taken by Rex Cardan, Ph.D. on 3/10/2014 using the Extradin A12 (XA132053) - RMC chamber.
Note on BU7: Correction measurements taken by Rex Cardan, Ph.D. on 3/10/2014 using the Extradin A12 (XA132053) - RMC chamber.
Note on BV7: Correction measurements taken by Rex Cardan, Ph.D. on 3/10/2014 using the Extradin A12 (XA132053) - RMC chamber.
Note on BW7: Correction measurements taken by Rex Cardan, Ph.D. on 3/10/2014 using the Extradin A12 (XA132053) - RMC chamber.
Note on BX7: Correction measurements taken by Rex Cardan, Ph.D. on 3/10/2014 using the Extradin A12 (XA132053) - RMC chamber.

## Physicists
The purpose of this tab is to maintain a list of physicists which can be selected on the "Measure" tab. If a new name is added, it will be available via the dropdown box on the "Measure" tab.
PHYSICISTS
Name
Rex Cardan, Ph.D.
Richard Popple, Ph.D.
Prem Pareek, Ph.D.
Ivan Brezovich, Ph.D.
Jun Duan, Ph.D.
Sui Shen, Ph.D.
Xingen, Wu, Ph.D.

## Corrections
The point of this page is to provide a tool for calculating and storing correction values Pion and Ppol during the annual QA. When you click the button at the bottom "PUSH TO SELECTED…." it stores all of these correction values for monthly use for the year. Specifically it stores them on the "Machines" tab under the selected accelerator from the "Measure" tab. A Pgrad tool is included but shifts are performed to the chamber in this worksheet so Pgrad is not actually used.
CORRECTION FACTOR WORKSHEET (MACHINE AND CHAMBER SPECIFIC)
Pion |  |  |  |  | Ppol
6X |  |  |  |  | 6X
Bias | Read 1 | Read 2 | Average |  | Bias | Read 1 | Read 2 | Average
150 | 26.65 | 26.65 | 26.65 | 1.0038 | -300 | 26.75 | 26.75 | 26.75 | 1.0011
300 | 26.75 | 26.75 | 26.75 |  | 300 | 26.81 | 26.81 | 26.81
15X |  |  |  |  | 15X
Bias | Read 1 | Read 2 | Average |  | Bias | Read 1 | Read 2 | Average
150 | 30.2 | 30.2 | 30.2 | 1.0063 | -300 | 30.39 | 30.39 | 30.39 | 1.001
300 | 30.39 | 30.39 | 30.39 |  | 300 | 30.45 | 30.45 | 30.45
6E |  |  |  |  | 6E
Bias | Read 1 | Read 2 | Average |  | Bias | Read 1 | Read 2 | Average
150 | 42.29 | 42.28 | 42.285 | 1.0142 | -300 | 42.87 | 42.88 | 42.875 | 1.0003
300 | 42.87 | 42.88 | 42.875 |  | 300 | 42.9 | 42.9 | 42.9
9E |  |  |  |  | 9E
Bias | Read 1 | Read 2 | Average |  | Bias | Read 1 | Read 2 | Average
150 | 42.59 | 42.57 | 42.58 | 1.0148 | -300 | 43.2 | 43.2 | 43.2 | 1.0005
300 | 43.2 | 43.2 | 43.2 |  | 300 | 43.25 | 43.24 | 43.245
12E |  |  |  |  | 12E
Bias | Read 1 | Read 2 | Average |  | Bias | Read 1 | Read 2 | Average
150 | 41.05 | 41.05 | 41.05 | 1.0141 | -300 | 41.61 | 41.63 | 41.62 | 1.0006
300 | 41.61 | 41.63 | 41.62 |  | 300 | 41.67 | 41.67 | 41.67
15E |  |  |  |  | 15E
Bias | Read 1 | Read 2 | Average |  | Bias | Read 1 | Read 2 | Average
150 | 40.04 | 40.03 | 40.035 | 1.0135 | -300 | 40.57 | 40.57 | 40.57 | 1.0012
300 | 40.57 | 40.57 | 40.57 |  | 300 | 40.7 | 40.64 | 40.67
18E |  |  |  |  | 18E
Bias | Read 1 | Read 2 | Average |  | Bias | Read 1 | Read 2 | Average
150 | 39.11 | 39.1 | 39.105 | 1.0132 | -300 | 39.61 | 39.62 | 39.615 | 1.0004
300 | 39.61 | 39.62 | 39.615 |  | 300 | 39.69 | 39.6 | 39.645
Note on F3: Rex Cardan, Ph.D.: Assumption is that you will use the -300V for measurements. Invert measurements if positive voltage is used.
PUSH TO SELECTED ACCELERATOR (FROM MEASURE TAB)
Clear Values
Clear Values

## CurrentSelections (hidden)
Selected Chamber |  |  |  | Selection Index
kQ | Low X | 0.99528 |  | 1
| High X | 0.98028 |  | 2
| kecal | 0.906 |  | 3
| rcav (cm) | 0.305 |  | 4
| Ndw-Co60 | 48270000 |  | 5
k'R50 | 6E | 1.0284 |  | 6
| 9E | 1.0171 |  | 7
| 12E | 1.0087 |  | 8
| 15E | 1.0033 |  | 9
| 18E | 0.99945 |  | 10
Calibration Date |  | 2016-11-09 |  | 11
Selected Electrometer |  |  |  | 12
Factor | 0.999 |  |  | 13
Calibration Date | 2015-11-05 |  |  | 14
15
BASIC INFORMATION | Machine Label |  | 21EX 1833 HSF AR | 16
| ENERGIES SUPPORTED | Low X | 6 | 17
|  | High X | 15 | 18
|  | E1 | 6 | 19
|  | E2 | 9 | 20
|  | E3 | 12 | 21
|  | E4 | 15 | 22
|  | E5 | 18 | 23
PHOTON INFORMATION | %dd(10)x | Low X | 66.6 | 24
|  | High X | 76.4 | 25
| Clinical PDD(10) | Low X | 0.668 | 26
|  | High X | 0.766 | 27
| Clinical TMR¦PDD(10) | Low X | 0.777 | 28
|  | High X | 0.87 | 29
| ISF | Inv. Sq. | 1.21 | 30
| dmax (cm, SSD,10x10) | Low X | 1.6 | 31
|  | High X | 2.7 | 32
ELECTRON INFORMATION | I50 (cm) | E1 | 2.364 | 33
|  | E2 | 3.547 | 34
|  | E3 | 4.955 | 35
|  | E4 | 6.246 | 36
|  | E5 | 7.562 | 37
| dref (cm) | E1 | 1.3235 | 38
|  | E2 | 2.0539 | 39
|  | E3 | 2.9232 | 40
|  | E4 | 3.7203 | 41
|  | E5 | 4.5328 | 42
| Electron Depth Dose at dref | E1 | 0.9992 | 43
|  | E2 | 0.9994 | 44
|  | E3 | 0.9987 | 45
|  | E4 | 0.994 | 46
|  | E5 | 0.9695 | 47
| R50 (cm, A10/A20 [for high E]) | E1 | 2.3 | 48
|  | E2 | 3.6 | 49
|  | E3 | 5 | 50
|  | E4 | 6.3 | 51
|  | E5 | 7.6 | 52
| R80 (cm, A10/A20 [for high E]) | E1 | 2 | 53
|  | E2 | 3 | 54
|  | E3 | 4.3 | 55
|  | E4 | 5.4 | 56
|  | E5 | 6.4 | 57
| R90 (cm, A10/A20 [for high E]) | E1 | 1.8 | 58
|  | E2 | 2.8 | 59
|  | E3 | 3.9 | 60
|  | E4 | 4.9 | 61
|  | E5 | 5.7 | 62
| Energy Tolerances (R80-2mm) | E1 | 0.9 | 63
|  | E2 | 0.9 | 64
|  | E3 | 0.85 | 65
|  | E4 | 0.84 | 66
|  | E5 | 0.82857 | 67
| Energy Tolerances (R80+2mm) | E1 | 0.6 | 68
|  | E2 | 0.7 | 69
|  | E3 | 0.71429 | 70
|  | E4 | 0.73333 | 71
|  | E5 | 0.75 | 72
CORRECTION INFORMATION | Cone Factor | E1 | 1 | 73
|  | E2 | 1 | 74
|  | E3 | 1 | 75
|  | E4 | 1
|  | E5 | 1
| Pion | Low X | 1.0039
|  | High X | 1.0058
|  | E1 | 1.015
|  | E2 | 1.0138
|  | E3 | 1.0118
|  | E4 | 1.0129
|  | E5 | 1.0151
| Ppol | Low X | 0.99981
|  | High X | 0.99966
|  | E1 | 1.0009
|  | E2 | 1.0002
|  | E3 | 1
|  | E4 | 1.0003
|  | E5 | 1.0007
SAD Type Calibration |  |  | TRUE
Note on B27: Rex Cardan, Ph.D.: This is used to check the energy ratio
Note on B29: Rex Cardan, Ph.D.: This is the clinical TMR(10,10x10) or PDD (10,10x10) that is used to move the dose from 10cm depth to dmax. If the machine is calibrated for 1cHy/MU at dmax, this value should be the TMR (10), otherwise it should be the PDD(10).
Note on B31: Rex Cardan, Ph.D.: This exists to correct for setup differences between SSD and Isocenter. If it is a TMR calibrated machine with 100 cm SSD setup, the ISF would be (110/100)^2. If it is a SSD calibrated machine, ISF = 1 and PDD(10) should be used in the TMR column.
Note on B64: Rex Cardan, Ph.D.: These values are interpolated between the R80 and R90 values.
Note on B69: Rex Cardan, Ph.D.: These values are interpolated between the R80 and R50 values.
Note on B74: Rex Cardan, Ph.D.: This is in place to adjust the electron measurements in the case the A10 cone is not used. If the A10 is used, just put in 1.0