## Autodrive
Energy | Mode | MU | Accessory | X1 | X2 | Y1 | Y2 | RepRate | Depth mm | Bias | N | M1 | M2 | M3 | Alert
6X | FIXED | 100 |  | 5 | 5 | 5 | 5 | 600 | 100 | 150 | 2 | -1.532e-07 | -1.529e-07 |  | Make sure SSD is 90cm
6X | FIXED | 100 |  | 5 | 5 | 5 | 5 | 600 | 100 | 300 | 2 | -1.535e-07 | -1.535e-07
6X | FIXED | 100 |  | 5 | 5 | 5 | 5 | 600 | 100 | -300 | 2 | 1.539e-07 | 1.538e-07
15X | FIXED | 100 |  | 5 | 5 | 5 | 5 | 600 | 100 | 150 | 2 | -1.771e-07 | -1.771e-07
15X | FIXED | 100 |  | 5 | 5 | 5 | 5 | 600 | 100 | 300 | 2 | -1.784e-07 | -1.783e-07
15X | FIXED | 100 |  | 5 | 5 | 5 | 5 | 600 | 100 | -300 | 2 | 1.787e-07 | 1.787e-07
6E | FIXED | 100 | A10 |  |  |  |  | 600 | 14.8 | 150 | 2 | -2.096e-07 | -2.098e-07 |  | Make sure SSD is 100 cm
6E | FIXED | 100 | A10 |  |  |  |  | 600 | 14.8 | 300 | 2 | -2.127e-07 | -2.127e-07
6E | FIXED | 100 | A10 |  |  |  |  | 600 | 14.8 | -300 | 2 | 2.126e-07 | 2.127e-07
9E | FIXED | 100 | A10 |  |  |  |  | 600 | 22.1 | 150 | 2 | -2.118e-07 | -2.12e-07
9E | FIXED | 100 | A10 |  |  |  |  | 600 | 22.1 | 300 | 2 | -2.146e-07 | -2.146e-07
9E | FIXED | 100 | A10 |  |  |  |  | 600 | 22.1 | -300 | 2 | 2.147e-07 | 2.148e-07
12E | FIXED | 100 | A10 |  |  |  |  | 600 | 30.8 | 150 | 2 | -2.143e-07 | -2.143e-07
12E | FIXED | 100 | A10 |  |  |  |  | 600 | 30.8 | 300 | 2 | -2.169e-07 | -2.171e-07
12E | FIXED | 100 | A10 |  |  |  |  | 600 | 30.8 | -300 | 2 | 2.173e-07 | 2.173e-07
15E | FIXED | 100 | A10 |  |  |  |  | 600 | 38.7 | 150 | 2 | -2.144e-07 | -2.146e-07
15E | FIXED | 100 | A10 |  |  |  |  | 600 | 38.7 | 300 | 2 | -2.171e-07 | -2.173e-07
15E | FIXED | 100 | A10 |  |  |  |  | 600 | 38.7 | -300 | 2 | 2.174e-07 | 2.178e-07
18E | FIXED | 100 | A10 |  |  |  |  | 600 | 46.9 | 150 | 2 | -2.12e-07 | -2.121e-07
18E | FIXED | 100 | A10 |  |  |  |  | 600 | 46.9 | 300 | 2 | -2.143e-07 | -2.146e-07
18E | FIXED | 100 | A10 |  |  |  |  | 600 | 46.9 | -300 | 2 | 2.148e-07 | 2.148e-07

## Corrections
The point of this page is to provide a tool for calculating and storing correction values Pion and Ppol during the annual QA. When you click the button at the bottom "PUSH TO SELECTED…." it stores all of these correction values for monthly use for the year. Specifically it stores them on the "Machines" tab under the selected accelerator from the "Measure" tab. A Pgrad tool is included but shifts are performed to the chamber in this worksheet so Pgrad is not actually used.
CORRECTION FACTOR WORKSHEET (MACHINE AND CHAMBER SPECIFIC)
Pion |  |  |  |  | Ppol
6X |  |  |  |  | 6X
Bias | Read 1 | Read 2 | Average |  | Bias | Read 1 | Read 2 | Average
150 | -1.532e-07 | -1.529e-07 | -1.5305e-07 | 1.0029 | 300 | -1.535e-07 | -1.535e-07 | -1.535e-07 | 1.0011
300 | -1.535e-07 | -1.535e-07 | -1.535e-07 |  | -300 | 1.539e-07 | 1.538e-07 | 1.5385e-07
15X |  |  |  |  | 15X
Bias | Read 1 | Read 2 | Average |  | Bias | Read 1 | Read 2 | Average
150 | -1.771e-07 | -1.771e-07 | -1.771e-07 | 1.0071 | 300 | -1.784e-07 | -1.783e-07 | -1.7835e-07 | 1.001
300 | -1.784e-07 | -1.783e-07 | -1.7835e-07 |  | -300 | 1.787e-07 | 1.787e-07 | 1.787e-07
6E |  |  |  |  | 6E
Bias | Read 1 | Read 2 | Average |  | Bias | Read 1 | Read 2 | Average
150 | -2.096e-07 | -2.098e-07 | -2.097e-07 | 1.0145 | 300 | -2.127e-07 | -2.127e-07 | -2.127e-07 | 0.99988
300 | -2.127e-07 | -2.127e-07 | -2.127e-07 |  | -300 | 2.126e-07 | 2.127e-07 | 2.1265e-07
9E |  |  |  |  | 9E
Bias | Read 1 | Read 2 | Average |  | Bias | Read 1 | Read 2 | Average
150 | -2.118e-07 | -2.12e-07 | -2.119e-07 | 1.0129 | 300 | -2.146e-07 | 0 | -1.073e-07 | 1.0005
300 | -2.146e-07 | -2.146e-07 | -2.146e-07 |  | -300 | 2.148e-07 | 0 | 1.074e-07
12E |  |  |  |  | 12E
Bias | Read 1 | Read 2 | Average |  | Bias | Read 1 | Read 2 | Average
150 | -2.143e-07 | -2.143e-07 | -2.143e-07 | 1.0128 | 300 | -2.169e-07 | -2.171e-07 | -2.17e-07 | 1.0007
300 | -2.169e-07 | -2.171e-07 | -2.17e-07 |  | -300 | 2.173e-07 | 2.173e-07 | 2.173e-07
15E |  |  |  |  | 15E
Bias | Read 1 | Read 2 | Average |  | Bias | Read 1 | Read 2 | Average
150 | -2.144e-07 | -2.146e-07 | -2.145e-07 | 1.0127 | 300 | -2.171e-07 | -2.173e-07 | -2.172e-07 | 1.0009
300 | -2.171e-07 | -2.173e-07 | -2.172e-07 |  | -300 | 2.174e-07 | 2.178e-07 | 2.176e-07
18E |  |  |  |  | 18E
Bias | Read 1 | Read 2 | Average |  | Bias | Read 1 | Read 2 | Average
150 | -2.12e-07 | -2.121e-07 | -2.1205e-07 | 1.0114 | 300 | -2.143e-07 | -2.146e-07 | -2.1445e-07 | 1.0008
300 | -2.143e-07 | -2.146e-07 | -2.1445e-07 |  | -300 | 2.148e-07 | 2.148e-07 | 2.148e-07
Note on F3: Rex Cardan, Ph.D.: Assumption is that you will use the 300V for measurements. Invert measurements if positive voltage is used.

## MEASURE
TG 51 Annual Calibration ¦ | [5 empty] | iX 5867 RMC
|  |  |  | Physicist | Rex Cardan, Ph.D. |  | Number of MU Used | 100
|  |  |  | Date |  |  | Temperature | 21.5 | C
|  | Ion Chamber (Cal: 04/17/2017)¦ Ndw :4.919E+7 |  |  | Extradin A12 (XA132053) - RMC |  | Pressure | 748.1 | mmHg
|  | Electrometer (Cal: 11/05/2015)¦ Pelec : 0.999 |  |  | SI Max 4000 PLUS - AR |  | Ctp | 1.0142
Photons (Bias +300V) ¦ 90SSD
6X
| Mraw (nC) |  |  | M (nC) | Dose Conversion |  | Energy Check | Output (cGy/MU) | OK?
Depth (cm) | Read 1 | Read 2 | Average | Mraw*Ppol*Pion*Ctp*Pelec | *Ndw - Co60 | *kQ |  | #REF!
10 | 15.35 | 15.35 | 15.35 | 15.6158 | 76.8143 | 76.4978 | 0.99477 | 0.995 cGy/MU @ dmax | PASSED
15X
| Mraw (nC) |  |  | M (nC) | Dose Conversion |  | Energy Check | Output (cGy/MU) | OK?
Depth (cm) | Read 1 | Read 2 | Average | Mraw*Ppol*Pion*Ctp*Pelec | *Ndw - Co60 | *kQ |  | #REF!
10 | 17.84 | 17.83 | 17.835 | 18.2162 | 89.6056 | 87.5951
Electrons (Bias +300V) ¦ 100SSD
6E
| Mraw (nC) |  |  | M (nC) | Dose Conversion (kR50) |  |  | Output (cGy/MU) | OK?
Depth (cm) | Read 1 | Read 2 | Average | Mraw*Ppol*Pion*Ctp*Pelec | *Ndw - Co60 | *kecal*k'R50/CF |  | M(dref) /PDD(dref)
1.15 | 21.27 | 21.27 | 21.27 | 21.8604 | 107.5312
9E
| Mraw (nC) |  |  | M (nC) | Dose Conversion (kR50) |  |  | Output (cGy/MU) | OK?
Depth (cm) | Read 1 | Read 2 | Average | Mraw*Ppol*Pion*Ctp*Pelec | *Ndw - Co60 | *kecal*k'R50/CF |  | M(dref) /PDD(dref)
1.15 | 21.46 | 21.46 | 21.46 | 22.0336 | 108.383
12E
| Mraw (nC) |  |  | M (nC) | Dose Conversion (kR50) |  |  | Output (cGy/MU) | OK?
Depth (cm) | Read 1 | Read 2 | Average | Mraw*Ppol*Pion*Ctp*Pelec | *Ndw - Co60 | *kecal*k'R50/CF |  | M(dref) /PDD(dref)
1.15 | 21.69 | 21.71 | 21.7 | 22.2818 | 109.604
15E
| Mraw (nC) |  |  | M (nC) | Dose Conversion (kR50) |  |  | Output (cGy/MU) | OK?
Depth (cm) | Read 1 | Read 2 | Average | Mraw*Ppol*Pion*Ctp*Pelec | *Ndw - Co60 | *kecal*k'R50/CF |  | M(dref) /PDD(dref)
1.15 | 21.71 | 21.73 | 21.72 | 22.3072 | 109.7289
18E
| Mraw (nC) |  |  | M (nC) | Dose Conversion (kR50) |  |  | Output (cGy/MU) | OK?
Depth (cm) | Read 1 | Read 2 | Average | Mraw*Ppol*Pion*Ctp*Pelec | *Ndw - Co60 | *kecal*k'R50/CF |  | M(dref) /PDD(dref)
1.13 | 21.43 | 21.46 | 21.445 | 21.9941 | 108.1892
Note on F12: Rex Cardan: The correction values are dependent on the machine and electrometer selected at the top of the page.
Note on H12: Rex Cardan: The value depends on the ion chamber selected. It is looked up via the "Chambers" tab
Note on I12: Rex Cardan: Rex Cardan: Looking at the current accelerator PDD (10) value from the "Machines" tab. Compares to within 2%
Note on B23: Rex Cardan, Ph.D.: This is the dref value + 0.5*rcav. This worsheet does not include pgrad corrections and instead picks measurement points that have the correction included (effective point of measurement is at dref)
Note on B27: Rex Cardan, Ph.D.: This is the dref value + 0.5*rcav. This worsheet does not include pgrad corrections and instead picks measurement points that have the correction included (effective point of measurement is at dref)
Note on B31: Rex Cardan, Ph.D.: This is the dref value + 0.5*rcav. This worsheet does not include pgrad corrections and instead picks measurement points that have the correction included (effective point of measurement is at dref)
Note on B35: Rex Cardan, Ph.D.: This is the dref value + 0.5*rcav. This worsheet does not include pgrad corrections and instead picks measurement points that have the correction included (effective point of measurement is at dref)
Note on B39: Rex Cardan, Ph.D.: This is the dref value + 0.5*rcav. This worsheet does not include pgrad corrections and instead picks measurement points that have the correction included (effective point of measurement is at dref)

## Chambers
CHAMBERS
Chamber Label | kQ Calc (from Table 1 - TG51) | [5 empty] | %dd(10)x |  | kQ |  |  |  |  | k'R50 |  |  |  |  | Last ADCL Calibration
| 58 | 63 | 66 | 71 | 81 | 93 | Low X | High X | Low X | High X | kecal | rcav (cm) | Ndw-Co60 | 6E | 9E | 12E | 15E | 18E
Extradin A12 (XA023335) - AR | 1 | 0.999 | 0.996 | 0.99 | 0.972 | 0.948 | 66.1 | 77.91 | 0.99588 | 0.97756 | 0.906 | 0.305 | 48270000 | #REF! | #REF! | #REF! | #REF! | #REF! | 2016-11-09 | 190
PTW TN30013 (1937) - RMC | 1 | 0.996 | 0.992 | 0.984 | 0.967 | 0.946 | 66.1 | 77.91 | 0.99184 | 0.97225 | 0.898 | 0.275 | 54960000 | #REF! | #REF! | #REF! | #REF! | #REF! | 2012-04-18 | 1856
Extradin A12 (XA132053) - RMC | 1 | 0.999 | 0.996 | 0.99 | 0.972 | 0.948 | 66.1 | 77.91 | 0.99588 | 0.97756 | 0.906 | 0.305 | 49190000 | #REF! | #REF! | #REF! | #REF! | #REF! | 2017-04-17 | 31
Extradin A12 (XA022324) - HSROC | 1 | 0.999 | 0.996 | 0.99 | 0.972 | 0.948 | 66.1 | 77.91 | 0.99588 | 0.97756 | 0.906 | 0.305 | 48040000 | #REF! | #REF! | #REF! | #REF! | #REF! | 2015-02-15 | 823
[7 empty] | 66.1 | 77.91 | #DIV/0! | #DIV/0! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 66.1 | 77.91 | #DIV/0! | #DIV/0! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 66.1 | 77.91 | #DIV/0! | #DIV/0! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 66.1 | 77.91 | #DIV/0! | #DIV/0! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 66.1 | 77.91 | #DIV/0! | #DIV/0! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 66.1 | 77.91 | #DIV/0! | #DIV/0! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 66.1 | 77.91 | #DIV/0! | #DIV/0! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 66.1 | 77.91 | #DIV/0! | #DIV/0! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 66.1 | 77.91 | #DIV/0! | #DIV/0! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 66.1 | 77.91 | #DIV/0! | #DIV/0! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 66.1 | 77.91 | #DIV/0! | #DIV/0! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 66.1 | 77.91 | #DIV/0! | #DIV/0! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 66.1 | 77.91 | #DIV/0! | #DIV/0! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 66.1 | 77.91 | #DIV/0! | #DIV/0! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF!
Note on J2: Rex Cardan, Ph.D.: These values will change based on the machine you choose on the Measure tab.

## Electrometers
The purpose of this tab is to keep a list of electrometers which can be used for the monthly QA. The date of the most recent ADCL calibration is also stored, so as to alert a user on the "Measure" tab when 1) It is a few months away from needed calibration and 2) Its calibration is expired
ELECTROMETERS
Electrometer Label | Electrometer Factor | Last ADCL Calibration
PTW 10005 (T111067¦High Mode) -AR | 1.003 | 2015-05-13 | 736
SI Max 4000 (E021141 ¦ High) - RMC | 0.999 | 2013-05-15 | 1464
SI Max 4000 PLUS - AR | 0.999 | 2015-11-05 | 560

## Machines
MACHINES MACHINES MACHINES
BASIC INFORMATION | [7 empty] | PHOTON INFORMATION |  |  |  | ELECTRON INFORMATION
| ENERGIES SUPPORTED | [6 empty] | %dd(10)x |  | Clinical TMR |  | I50 (cm) |  |  |  |  | dref (cm) |  |  |  |  | Electron Depth Dose at dref |  |  |  |  | Cone Factor
Machine Label | Low X | High X | E1 | E2 | E3 | E4 | E5 | Low X | High X | Low X | High X | E1 | E2 | E3 | E4 | E5 | E1 | E2 | E3 | E4 | E5 | E1 | E2 | E3 | E4 | E5 | E1 | E2 | E3 | E4 | E5
21EX 1833 HSF AR | 6 | 15 | 6 | 9 | 12 | 15 | 18 | 66.6 | 76.4 | 0.777 | 0.87 | 2.364 | 3.547 | 4.955 | 6.246 | 7.562 | 1.3235 | 2.0539 | 2.9232 | 3.7203 | 4.5328 | 0.9992 | 0.9994 | 0.9987 | 0.994 | 0.9695 | 1 | 1 | 1 | 1 | 1
iX HSF 5326 AR | 6 | 15 | 6 | 9 | 12 | 15 | 18 | 66.15 | 77.6 | 0.777 | 0.87 | 2.312 | 3.492 | 4.94 | 6.175 | 7.39 | 1.2914 | 2.02 | 2.914 | 3.6764 | 4.4266 | 0.9991 | 0.999 | 0.9984 | 0.9944 | 0.975 | 1 | 1 | 1 | 1 | 1
iX 5867 RMC | 6 | 15 | 6 | 9 | 12 | 15 | 18 | 66.1 | 77.91 | 0.769 | 0.877 | 2.396 | 3.63 | 5.076 | 6.438 | 7.845 | 1.3433 | 2.1052 | 2.9979 | 3.8388 | 4.7075 | 0.999 | 0.999 | 1 | 0.994 | 0.974 | 1 | 1 | 1 | 1 | 1
Note on K3: Rex Cardan, Ph.D.: This is the clinical TMR(10,10x10) or PDD (10,10x10) that is used to move the dose from 10cm depth to dmax. If the machine is calibrated for 1cHy/MU at dmax, this value should be the TMR (10), otherwise it should be the PDD(10).
Note on AB3: Rex Cardan, Ph.D.: This is in place to adjust the electron measurements in the case the A10 cone is not used. If the A10 is used, just put in 1.0

## Physicists
The purpose of this tab is to maintain a list of physicists which can be selected on the "Measure" tab. If a new name is added, it will be available via the dropdown box on the "Measure" tab.
PHYSICISTS
Name
Rex Cardan, Ph.D.
Richard Popple, Ph.D.
Prem Pareek, Ph.D.
Ivan Brezovich, Ph.D.
Jun Duan, Ph.D.
Sui Shen, Ph.D.
Xingen, Wu, Ph.D.

## CurrentSelections
Selected Chamber |  |  |  | Selection Index
kQ | Low X | 0.99588 |  | 1
| High X | 0.97756 |  | 2
| kecal | 0.906 |  | 3
| rcav (cm) | 0.305 |  | 4
| Ndw-Co60 | 49190000 |  | 5
k'R50 | 6E | #REF! |  | 6
| 9E | #REF! |  | 7
| 12E | #REF! |  | 8
| 15E | #REF! |  | 9
| 18E | #REF! |  | 10
Calibration Date |  | 2017-04-17 |  | 11
Selected Electrometer |  |  |  | 12
Factor | 0.999 |  |  | 13
Calibration Date | 2015-11-05 |  |  | 14
15
BASIC INFORMATION | Machine Label |  | iX 5867 RMC | 16
| ENERGIES SUPPORTED | Low X | 6 | 17
|  | High X | 15 | 18
|  | E1 | 6 | 19
|  | E2 | 9 | 20
|  | E3 | 12 | 21
|  | E4 | 15 | 22
|  | E5 | 18 | 23
PHOTON INFORMATION | %dd(10)x | Low X | 66.1 | 24
|  | High X | 77.91 | 25
| Clinical TMR | Low X | 0.769 | 26
|  | High X | 0.877 | 27
ELECTRON INFORMATION | I50 (cm) | E1 | 1.3433 | 28
|  | E2 | 2.1052 | 29
|  | E3 | 2.9979 | 30
|  | E4 | 3.8388 | 31
|  | E5 | 4.7075 | 32
| dref (cm) | E1 | 0.999 | 33
|  | E2 | 0.999 | 34
|  | E3 | 1 | 35
|  | E4 | 0.994 | 36
|  | E5 | 0.974 | 37
| Electron Depth Dose at dref | E1 | 1 | 38
|  | E2 | 1 | 39
|  | E3 | 1 | 40
|  | E4 | 1 | 41
|  | E5 | 1 | 42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65
66
67
68
69
70
71
72
73
74
75
Note on B27: Rex Cardan, Ph.D.: This is the clinical TMR(10,10x10) or PDD (10,10x10) that is used to move the dose from 10cm depth to dmax. If the machine is calibrated for 1cHy/MU at dmax, this value should be the TMR (10), otherwise it should be the PDD(10).